## Hazard Thresholds
Hazard | Threshold for consideration | Review Guidance
Mechanical Hazards | Weight exceeds 3 tons and is supported above the floor | Mechanical Structures and Systems - The design of mechanical structures and systems shall be evaluated to verify:
| Exceeds 10 tons in total weight | 1.Calculations have been prepared and are based on good engineering practice.
| Moves at a speed greater than 5 ft/sec | 2. Calculations reference and follow applicable standards and codes including but not limited to ASME Section VIII, ANSI, AWS, AISC 9th Edition ASD, CGA, NFPA as appropriate.
| Costs more than $100k to replace | 3. Calculations address each element of the design in sufficient detail to ensure the complete structure or system meets applicable design standards and that the completed structure or system minimizes occupational safety hazards. The reviews should emphasize the overall design process rather than simply checking what has been provided to minimize oversights as part of the design process.
| Includes pressure/vacuum vessels | 4. Assembly procedures for elements are prepared as deemed necessary by the NuMI ES&H/QA Review Committee and/or NuMI ES&H Coordinator in order to minimize occupational safety hazards during installation or assembly and to minimize delays during installation
Flammable Gas | Presence of flammable gases | Flammable gas systems shall be reviewed for compliance with FESHM 6020.3.
Electrical | Any Primary Beam Magnet Power Supply Having Output Capability Greater Than 50 Amps | Review to be conducted and documented by Beams EE Support Department in accord with Departmental Procedure BDDP-EE-1002. Such a review typically involves generation of a Safety Analysis Report, safety walkthrus before commissioning and/or operation, and determination as to necessity of and development of written LOTO procedures.
| Kicker Power Supply System | Review to be conducted and documented by Beams EE Support Department in accord with Departmental Procedure BDDP-EE-1002. This review involves generation of a Safety Analysis Report, a safety walkthru before commissioning and/or operation, and development of the associated written LOTO procedure.
| Horn Power Supply System | Review to be conducted and documented by Beams EE Support Department in accord with Departmental Procedure BDDP-EE-1002. This review involves generation of a Safety Analysis Report, a safety walkthru before commissioning and/or operation, and development of the associated written LOTO procedure.
| Hadron and Muon Monitors & MINOS Near Detector Experimental Apparatus | Review to be conducted in accord with established Particle Physics Division procedures for review of experiments before operation. Special precautions may be required in accord with particularly high density/value of associated electronics. Special consideration shall be given to non-commercial, low-voltage, high-current distribution systems and associated requirements of FESHM Chapter 5046.
Fire Hazards | Large inventory of combustible items | Responsibility for the fire protection program are addressed in FESHM 6010. The responsibility of the NuMI ES&H/QA Review Committee is to verify compliance with FESHM 6040.1 and 6040.2 for items included as part of the project. The intent of this review is to verify compliance during the design phase to avoid remedial action and resulting additional costs or installation delays.
Oxygen Deficiency Hazards | Presence of oxygen displacing gases | Systems that may result in oxygen deficiency hazards shall be reviewed for compliance with FESHM 5064.
Magnetic Hazards | Presence of magnetic fields | Systems that have static magnetic fields shall be reviewed for compliance with FESHM 5062.2
Cryogenic Hazards | Presence of cryogenic liquids | Systems that may result in oxygen deficiency hazards shall be reviewed for compliance with FESHM 5032.
Laser Hazards | List any lasers that will be used | Systems, which utilize lasers, shall be reviewed for compliance with FESHM 5062.1.
Radiation Hazards | Activated components exceeding 1 R/hr at 1 foot | The responsibility of the NuMI ES&H/QA Review Committee is to verify that the ES&H Section addresses radiation safety. Radiation safety reviews shall satisfy the Fermilab Radiological Control Manual.
List radioactive sources
Toxic Materials | Inventory of toxic or hazardous materials | This review shall verify that elements are designed and built to minimize occupational exposure to toxic materials.
Occupational Industrial Hazards | Presence of occupational industrial hazards | The responsibility of the NuMI ES&H/QA Review Committee is to verify that the Division’s ES&H Department addresses occupational industrial hazards.
Any/All | Use of computers in systems that protect people, property or the environment | Computers and Programmable Logic Controllers (PLC's). The lab requires that detectors or apparatus that rely solely on dedicated computers or PLC's for safety, environment, or property protection functions must comply with Directors Policy # 21 (see http://www.fnal.gov/directorate/Policy_Manual.html#No_21).